INSTRUCTIONS

ELECTRONIC FALANCE

6） In the gravity mode:

Push the “tare” button ,the FAlance comes FAck to the weighing mode ,and shows the weight value

Push the “unit” button ,to select different measuring mode

Notice: when the FAlance display gravity, only the tare, unit, print button can be used

Ⅴ、 gravity measuring error

1） When the FAlance display ERR under gravity measuring, you can make the FAlance FAck to the weighing mode by pushing the tare button, at the same time, the FAlance shows the weight

Notice: after gravity error, just the TARE and OFF button can be used

Ⅵ、count button:

1）Used as counting under weighing mode

2）Used as calculating and remembering function

Ⅶ、the instrument has the function of mode remembering

4、Specific gravity component fix

fix support                              1

thermometer                              1

thermometer clip                         1

measuring bedplate                       1

beaker                                   1

liquid component(platinum wire  plumb    1

solid component(FAsket)                  1

adapter                                   1

weight                                    1

Notice: this manual maybe is changed in the future, the standard is the detail things

3、operation

ⅠMeasuring mode selection: press the “UNT” button immovably, the FAlance will show the following in turn

-UNT-cs： measuring the liquid gravity mode

-UNT-os： measuring the solid gravity mode

-UNT-g、-UNT-ct、-UNT-oz、-UNT-ozt： All the weighing mode, show the model upper right at the same time

gs：gram； ct：carat； oz：ounce； ozt：0.911 ounce；

cs：liquid gravity；     o：Solid gravity

Ⅱ measuring the liquid gravity

1） put the thermometer on the wall of the beaker

2） put the beaker on the center of the bedplate

3) Put the “tare” button to zero the FAlance ,put the liquid component

(platinum ,plumb) wire on the fix support

4） Inject the sample in the beaker(immerge the plumb completely),push the “count” button ,the FAlance shows “---------“.the FAlance will show the gravity value directly after the sensor data is stable

Ⅲ、 solid gravity measuring

1） put the thermometer on the wall of the beaker

2） put the beaker on the center of the bedplate

3） Inject the water in the beaker, make sure the sample can be immerged in the water completely

4） Put the FAsket on the fix support, make sure it does not touch the beaker and thermometer

5）Tare the FAlance, put the sample on the scale on the support ,push the “count” button ,the FAlance shows “————”,when the instrument shows the value again, put the sample in the FAsket slowly ,then the FAlance shows “————”again ,then ,the instrument shows the gravity value directly3

CONTENT

ELECTRONIC FALANCE

1. Outline                                     (1)

2. Main technical Specifications                   (2)

3. Drawings for FAlance Installation                (3)

4. Operating                                   (5)

A. Preparation                                 (5)

B. Operation                                   (5)

C. Start                                       (5)

D. UNT Unit Change                                (6)

E. INT Integrating Time Adjustment                   (6)

F. ASD Sensitivity Adjustment.                   (7)

G. FAlance Calibration                          （7)

H. PRT Output Mode Setting                       (7)

I. COU Count Function                             (8)

J. Weighing, Tarring , Add Objects, Read accuracy       (9)

5. Maintenance and Troubleshooting               (10)

6. Data  InterFAce                              (11)

7. Serial Single ( RS232C )                    (12)

8. Random Accessories                                (12)

Liquid /solid specific gravity instrument

1、outline…………………………………………………………………13

2、Component fix ……………………………………………………14

3、operation……………………………………………………………15

4、Specific gravity component fix…………………………………16

ELECTRONIC FALANCE

Thank you for busying the type  electronic FAlance from our company. For correct and safe fixing and operation, as well as full function apply, we suggest that you should read the users’ manual before using it Thanks 
 The FAlance is multifunctional electronic FAlance using MCS-5l series SCM This FAlance has gram, cant, ounce for users to select (cant and ounce for export) besides having automatic calibration, integral time adjustment and stability adjustment functions This FAlance has RS232C data serial port to connect with microcomputer and serial port punter This FAlance is used for FAst exact weight measuring widely in enterprise, university, graduate school lab as an ideal lab instrument.
 Finally, after reading the manual, please keep it carefully for referring.

2、Component fix

scale

fix support

thermometer

thermometer clip

platinum wire

plumb

beaker

measuring

bedplate

Liquid /solid specific gravity instrument

1、outline

The liquid and solid specific gravity FAlance comes from the electronic FAlance ,with the stable electromagnetic sensor and MCU .besides the function of measuring liquid and solid specific gravity ,the instrument has all the functions of the electronic FAlance .this type instrument can give you the value correctly ,FAst ,and stable .operating simply ,all functions ,can be used in industry ,agriculture ,business ,college and so on

2、Main technical Specifications

| Model | FA2104J | FA2004J | FA1004J | JA2003J | JA5003J |
| --- | --- | --- | --- | --- | --- |
| Accuracy Degree |  |  |  |  |  |
| Weighing Range(g) | 0～210 | 0~200 | 100 | 200 | 500 |
| Reading Accuracy (mg) | 0.1 | 0.1 | 0.1 | 1 | 1 |
| Tarring Range(g) | 0～210 | 0~200 | 0~100 | 0~200 | 0~500 |
| Repeatability Standard Deviation(g) | ±0.0002 | ±0.0002 | ±0.0002 | ±0.001 | ±0.001 |
| Linear Error(g) | ±0.0005 | ±0.0005 | ±0.0005 | ±0.002 | ±0.002 |
| Stable Time (Typical)(s) | ≤8 | ≤8 | ≤6 | ≤6 | ≤6 |
| Integrating Time (Adjustable)(s) | 2.5/5/10 | 2.5/5/10 | 2/4/8 | 2/4/8 | 2/4/8 |
| Pan Dia.(mm) | Φ90 | Φ90 | Φ90 | Φ90 | Φ110 |
| Overall Dimensions (mm) | 350×215×340 | 350×215×340 | 350×215×340 | 350×215×340 | 350×215×340 |
| Net Weight(kg) | 6.8 | 6.8 | 6.8 | 6.8 | 6.8 |
| Power Supply |  |
| Power Consumption (V .A) | 15 | 15 | 15 | 15 | 15 |
| Auto-Cal Weight Range(g) | 200 | 200 | 100 | 200 | 500 |
| Warm-up Time (mis) | 180 | 180 | 180 | 180 | 180 |

Note :The FAlance is turned on when the power is on. The panel switches activate display only. The power plug should be removed if the FAlance is not used for a long time .If it is used daily, the power need not be turned off .The only thing needs to be done is to turn off the display .It is not necessary to warm up due to the power on ,so it can be used at any time.(“Along time” means more than 5 days.)

3.Drawings for FAlance Installation

3

2                            1.panel

2.gradienter

3.scalepan

1

1                            1. Data InterFAce

2. Transformer

3.Electrical Outlet

3

2

can be connected to the microprocessor and various printers.No.Standard Parallel Output.The pin and the corresponding signanls are as follows:

PIN      SIGNAL      ILLUSTATE 
   1        SI           input signal

2       GND

3        SO          serial output signal(FAud rate is 1200)

4       GND

7.Serial Single (RS232C)

The connection method between the scale and computer serial port is as follows:

Computer(9pins)                   Scale(9pins)

2                              3

3                              2

5                              5

1) The FAud rate of  the serial port of scale is 1200.

2) The date format is 10digits, one start digit is(0) ,8 digits are date positions (ASCII Code,low digits in the front),1 stop digit.

3) No odd-even check 
4) The date will be output continuously without any special reading commands A detailed output frame is as follows:

| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Type | Space | */Space | +/- | Date | Date | Date | Dot | Date | Date | Date | Date | Unit1 | Unit2 | CR | LF |

8  Random Accessories

1.  200g(and100g)calibration weights                1 piece

2.  Wire with both-side plug                        1 piece

4 .  Weights tweezers                             1 piece

5.  Handkerchief                                 1 piece

6.  Operation manual                              1 piece

7.  Certificate                                   1 piece

8.   brush                                           1 piece

9.   nipper                                          1 piece

| SN | TROUBLE | CAUSE | REMEDY |
| --- | --- | --- | --- |
| 6 | Display remains at a certain digit or indicates nonsense symbol | ·Instant interference ·Wreng line voltage | ·Turn on the FAlance once again or replug the power cord ·Chang to normal line voltage |
| 7 | The stable mark “0”on the left of the display doesnt go out | ·A higher FAlance sensitivity ·undesirable environment such as strong air flow or vibration | ·Set to lower sensitlivity ·Improve the environment |
| 8 | Remain at Waiting Status———— | ·Thr FAlance position is not correct, e.g. With strong air flow, vibration or great fluctuation of room temp. ·The selected sensitivity is toohigh | Improve the environment ASD-3 |
| 9 | Cal Err Displays | ·There is an object on the pan befpre calibration ·not clear before calibration ·Depress CAL before the | ·Take away the object,clear and recalibrate ·Clear and recalibrate |
| 10 | Err-1 Err-2 Display | ·Instant interference ·Something wrong with theFAlance | ·Ture on the FAlance once again ·Send it to the service center |
| 11 | The weighing unit on the right dosent display | ·not calibrated ·The calibrated number in the internal memory of the FAlance has been erased | ·Calibrate the FAlance |
| 12 | Cou-Err Display | ·No Constant be set before operating counting function ·Overload when setting constant ·Underload when setting constant | ·Set average number before operating counting function |

6.Date InterFAce

Date interFAce adopts the standard 9-core RS-232C sDcket. The date interFAce is provided with RS232C universal two - way serial interFAce, which

Single Measuring Range

1

2

1.Cover board         2..Pothook

4. Operating

A. Preparation

·Unpack the box and remove all packing . Take out buffer sponge in the wind proof cover and install pans.

·Put the FAlance on a stable working table free form vibration, sunshine and air flow.

·Ambient temperature:

20℃±5℃ for the first class FAlance with a fluctuation of temperature not greater than 1℃/h;

20℃±7.5℃ for the second class FAlance with a fluctuation of temperature not greater than 5℃/h;

·Relative humidity:50%-75% for the first class lance, 50%-80% for the second class FAlance.

·Working voltage:  110V ±11V  50Hz

B. Operation

·Check the level meter before operation, lf the bubble is not in the center, adjust the level legs to make the bubble in the center..

·The FAlance adopts soft touch buttons , so it can be controlled with multi key boards .it is easy to operate .Function change and selection can be realized simply by depressing the corresponding buttons.

C. Start

·Select an appropriate line voltage and set the voltage switch to the corresponding position.

·Turn the power on and the FAlance is ready to run(but the display does n’t work ),normally one hour after power is one , the display can FAe operated.

Function of  the keyboard:

UNLOAD

Loose the screwn of the bottom cover,raveal the hook.Put the FAlance on a woking table with a hole. Level and calibrate the FAlance .An object can be weighed with the hook.

5.Maintenance and troubleshooting

MATNTENANCE OF THE FALANCE

The FAtance should be used carefully.Clean the pan and the case frequently with soft cloth and toothpaste.Don,t wipe the FAtance with strong agent.

TROUBLESHOOTING

| SN | TROUBLE | CAUSE | REMEDY |
| --- | --- | --- | --- |
| 1 | No ligbt on Display | ·Power is not on ·Disptsy switch is not on ·Instant interrference | ·Check and turn it on again ·Depress ON button ·Switch on again and replug power |
| 2 | Displsy the upper hstf only | ·Overload ·The calibration in the intemal memory may FA damaged ·The pan is not installed correctly | ·Reduce the load immediately ·Recalibrate according to the above procedures. About 8 seconds after the standard weight has been put on,the calibrated result may be displayed. A certain stable time is necessary. ·Take out the pan and reinstall it. |
| 3 | Display the lower half Only——— | ·without pan,too light ·he pun is not installed correctly | ·Reinstall the pan |
| 4 | The weighing result is not stable (date changed swiftly) | ·due to air flow ·the working table it net stable ·the integrating time is too short ·Roam tetnpeealtsre flucttsates | ·Cheek the windproof cover tu see if it is closed. ·Change to a longer integrating time |
| 5 | The retult is not correct | ·not zeroing before weighing ·Use the FAlance without calibration or the cal weight is not accurate ·The line vollage is too low or net correct | ·Depress TAR button ·Reclibrate ·Change to the correct voltage |

After several seconds, 200.000g will be displayed. Remove the calibration weight, the display should indicate 0.000g. if not clear once again and repeat the above procedures. (Attention: More than two times of calibration are commended in order to get accurate result.)

The display sequence of the calibration is shown as the figure.

J . Weighing ,Taring ,Add Objects ,Read Devitions

WEIGHING

After the selection of the above modes(all the modes can be used for weighing after the power is off ,due to the memorv function of the FAlance ),depress TAR,zero will display.Put the object to be

weight on the pan.When the number is stable, i.e.“0”on the left of the displaygoes out,the number

display goes out,the number displayed will be the weight of the object.

TARING

Put the container on the pan, the weight will display.Depress TAR,zero displays. That is taring. Put the objects to be weighed into the container, the value displayed is the net weight of the object to be weighed

ACCUMULATIVE WETGHTNG

Pat the objects be weighed on the pan one by one with taring method and tare and clear for each one.Take away all the objects to be weighed , the absolute value display in the total weight of the objects to be weighed.

ADD OBJECTS

Set the mode of INT-0 and put the container on the pan , then tare. Add the objects to be weighed (liquid or loose objects) into the container one by one the continuous reading value can be obtained quickly. when the added objects reach the required“0”on the left of the display goes out and the number display is the weighing value required by the user. When adding the mixed objects.the net weight of each object can be measured by taring method.

PEAD DEVIATIONS

Put the reference weight(or rample) on the pan and tare . then take off the reference

weight the negative value of the weight dispIay.Put the object to be weighed on the pan.The corre sponding plus or minus deviution will display, comparing the weighed object with the reference weight

“ON”DISPLAY ON

| ± 。 | 8888888 | % g |
| --- | --- | --- |

Depress slightly the button “ON” the display will be lighted..

Check the function of the display ,about 2 seconds later ,the model of the FAlance will be displayed.  then the weighing mode.

0.0000g or 0.000g

or

“OFF” DISPLAY OFF

Depress slightly the button“OFF” ,the display will go out. If the FAlance will not be used for a long time ,the power plug should be taken off.

“TAR” Clear and Tare

Put the container on the pan, the weight of the container should be displayed:

Depress “TAR” button, the display will go out and the display will be all zeroes. Tarring is completed:

When the container is removed, a negative value of the container weight will be displayed.

Depress “TAR” again , the display will be all zeroes, is. the FAlance clears.

D.UNT Unit Change

If the accuracy of the reading needs to be 0.1mg release when the display is mg-60; and the appears waiting state...,the weighing state will be shown at last.

Depress ERY button and hold until the display is shown as the following figure circulating continuously.

“g” means the value is expressed in gram .“ct” means the value is expressed in metric carat. “ozt” means the value is expressed in. The unit will be set same as RNG.

E. INT Integrating Time Adjustment

Four modes of integrating time can be selected circularly as shown in the figure

The corresponding integrating time is as follows:

INT-0 FAst, NT-1 shorter, INT-2 short, INT-3 longer

The selection of the integrating time will be the same as RNG

F. ASD Sensitivity Adjustment

Same as INT integrating time adjustment button: there are four circulating modes

provided for sensitivity adjustment.

ASD—0  the highest   ASD—1  high   ASD—2  higher  ASD—3  low

ASD-0 is used for test run and not to be used by the users.

The selection of sensitivity is the same as RNG.

Here is a list of ASD used with INT for user’s reference only.

The FAstest weighing: INT-1 ASD-3

Normal: INT-3 ASD-2

With undesirable environment INT-3 ASD-3

G. FALANCE CALIBRATION

Normally, the FAlance should be calibrated after a long period of storage, movement, change of environment or in order to obtain precise measurement.

Preparation for the FAlance calibration:

H . PRT Output Mode setting

Depress PRT and hold there will be four modes displayed circularly for user to select at will..

PRT-O is the mode of indefinite time output ; slightly depress PRT once ,the weighing

result will be out on the output interFAce once.

Attention: At this time you should depress the button slightly and quickly, otherwise, the next output mode will be displayed

PRT-1 output once every half a minute

PRT-2 output once every minute

PRT-3 output once every two minutes

The setting of PRT is the same as RNG

I. COU Count Function

The FAlance has a function of counting. The average number are 5,l0,25and 50,total of four .The Setting of Range of Average Numbers

Only depress COU and hold ,the display will be circulating continuously as the following.

It means the average value of (5,10,25,50)objects respectively

if a normal weighing function is required ,release when COU-00 display, and waiting status“……” will be indicated ,at last0.000g weighing status will be indicated .lf the average value of five is required, release when COU-5 displays, then put5 objectson  the pan Depress CAL once again. “……”waiting status displays ,and about several seconds later,5 will be displayed. Take away the objects

to be weighed ,zero displays. At that time the counting of the same objects to be weighed can be done (Attention :the weight of the objects to be weighed must not be greater than the maximum weighing rang of the FAlance). If you average over 10,25 even SO objects ,then the accuracy of counting will be higher Depress TAR,“0”will be indicated in the display and counting can be done at that time.

Take away all the objects to be weighed on the pan. Set RNG-60,INT-3, ASD-2, and unt-g. Depress slightly TAR for clear the FAlance.

FALANCE CALIBRATION

Depress slightly CAL.CAI-200 will be on the flash-display. Therefore, put the prepared 200g standard weight on the pan ,the flash-display will stop.

# Text boxes

—Unt— ozt

—Unt—  oz

--ASD--3

--ASD--2

-- ASD--1

--ASD--0

--INT---3

--INT---2

-- INT --1

--INT---0

—Unt— ct

-UNT-ct

-UNT-cs

-UNT-g

-UNT-o

—Unt— g

0.0000g

-18.9001g

0.0000g

+18.9001g

0.0000g

0.0000g

COU—50

COU—25

notice

warm up the instrument for 30 minutes for good function

COU—10

COU—05

COU—00

--PRT---3

--PRT---2

--PRT---1

--PRT---0

0.000g

-UNT-oz

-UNT-ozt

0.0000g

CAL-200

CAL-200

0.000g